PreciseFlex™ Robot Force Based Height Detection

Version 3.0D1, September 2023

“HeightDetect” is a software package for PreciseFlex™ robots that permits the height of various objects, such as a stack of sample plates, to be detected by automatically monitoring the Sample Handler’s motor torques.  Robots such as the PreciseFlex 3400 are ideally suited for this task due to their geometry, low-friction back-drivable axes, and software gravity compensation.  With proper analysis, these features permit contact forces to be detected without the need for extra hardware sensors.  In contrast, 6-axis articulated robots typically are very heavy, have high friction axes, and are equipped with very high gear ratio motor drives that are difficult to back-drive.  These features make it impossible to detect small contact forces without additional sensors and even make it very difficult to delicately position objects.

Using this height detection function is very straightforward. Once the gripper is positioned above the object to be measured, the HeightDetect GPL routine is executed.  This function: moves the gripper down until a specified resistance force is detected; analyzes the collected data to compute the position of the robot when contact is first made and returns the contact position to the calling routine.  The results typically have an average repeatability of 0.5 mm and a worse case repeatability of 1.5 mm.  For best results, the gripper should be positioned at least 20mm above the expected contact surface, although this function will provide good results with separations of 10mm. The entire process typically takes less than 10 seconds to execute.

For even better results, a “thorough” secondary data collection and analysis can be performed immediately after the initial process as the gripper is being lifted up from the contact position.  This reduces the typical average repeatability to 0.3 mm and the worse case repeatability to 0.6 mm.  This secondary operation adds approximately 4 seconds to the entire process.

HeightDetect is part of a larger force based software package that permits the Sample Handler to automatically adjust to position errors and to automatically teach plate nest locations.  This function is provided as a software module that can be executed as part of the TCP Command Server or installed as part of a customer specific GPL software project.

### 1.   System Requirements

The requirements for the execution of this package are as follows:

- Any model of the PreciseFlex™ 3400 or DD4 family of Sample Handlers with various lengths of Z-travels, ranges of reaches, and grippers and fingers, or a PrecisePlace 100 with gripper and fingers.

- A copy of the TCP Command Server Version 2.1 or later that includes the Auto Center plug-in module.  For DD4 support, TCS 3.0D1 or later is required.

- GPL 3.2I5 or later installed on the controller for PreciseFlex 400

GPL 4.0G11 or later for the PrecisePlace 100.

GPL 5.1A2 or later for the PreciseFlex DD4.

- A “Precise Application” or “Z Height Detection” License installed on the controller for the Sample Handler.

### 2.   Executing HeightDetect Within the TCP Command Server

The HeightDetect function is provided as part of the “PARobot_auto_center” plug-in module to the TCP Command Server (TCS).   This plug-in module is pre-installed in the Precise robot version of the TCS, so this function is ready to run so long as the system satisfies the requirements stated above.

The following describes the TCS command syntax.

| Command Case ignored | Arguments Defaults values are set if no arguments are specified | Description, responses are for PC mode. Telnet mode is similar but more descriptive. |
| --- | --- | --- |
| HeightDetect | arg1: QuickMode, specifies if the function should be executed more quickly with less accuracy or more slowly and more accurately. 1 = Quick execution only. 0.5 mm typical and 1.5 mm worse case repeatability errors. 2 = Execute Quick mode followed by Thorough mode. 0.3 mm typical and 0.6 mm worse case repeatability errors. Adds about 4 seconds to total process time (default). arg2: rXSearchLimit, maximum distance that the gripper is permitted to move down relative to the starting position, in mm. Set to -1E6 by default. arg3: XMaxForce, maximum vertical force to be applied when searching for the stack height, in Nt. Set to -15 by default. | Executes the HeightDetect function that moves the gripper down and detects the height of a stack of plates or the height of a nest using motor force sensing. Prior to executing this command, the gripper should be positioned over and at least 10-20mm above the stack. At the completion of this command, the gripper is left a few mm above the located surface. The function was designed to execute with the system speed set to 100%. Slower system speeds can be utilized, but this speed should not be lower than 5%. Given the complexity of this process, it does not issue a reply until the function has completed. When returned, the reply contains the status code followed by the detected height in mm, measured from the robot’s World Z=0 plane. |

### 3.  Installing HeightDetect In A Custom GPL Application

HeightDetect can be executed from within the TCP Command Server project in which it is distributed or it can be included in a customer specific application.  To include this function within your own custom GPL project, do the following:

- Launch the Guidance Development Environment.

- In the Project Manager window, right click on your GPL project.

- Select “Add new item”.

- In the Add New Item pop-up, click on the “Code File” icon, fill in the Name field with “PARobot_auto_center.gpp” and press Add.  This will add a dummy placeholder code file into your project.

- Using Windows Explorer, locate your copy of the TCP Command Server project and copy the PARobot_auto_center.gpp” file to your own project folder.  Replace the dummy placeholder file with the real code file.

The HeightDetect function can now be executed from within your custom application.

### 3.1   Execution

The HeightDetect function consists of three top-level callable GPL routines:

AC_InitAutoCenter   -  This routine must be called once at the start of your code to pre-compute a number of constant values that are utilized by the HeightDetect function.

AC_ErrorMessage  -  This routine converts both standard system exceptions and custom exceptions generated by HeightDetect into a String for output to the operator.

AC_HeightDetect   -  This is the main entry point for the HeightDetect function.  It must be executed once each time you wish to measure the height of an object.

The calling sequence for each of these routines is provided in Appendix A.

The HeightDetect function was designed to be executed with the System Speed set to 100%.  Slower system speeds can be utilized, but this speed should not be lower than 5%.

#### 3.2    Sample Software

The following is a simple GPL program that illustrates how HeightDetect can be repeatedly executed to determine the height of an object.  The Location “start_j” defines a position above the object.

Public Sub MAIN

' ABSTRACT:  Main program to exercise the Height Detection function.

'

' INPUTS:	None

'

' OUTPUTS:	None

Dim start_j As New Location

Dim stack_top As New Location

Dim to_start As New Profile

Dim Rb As Integer = 1

Dim exc1 As New Exception

Const CRLF As String = Chr(GPL_CR) & Chr(GPL_LF)

start_j.Angles(335.382,62.337,230.45,66.974,78)	' Sample starting point

to_start.AccelRamp = 0.1				' Move to start with this speed

to_start.DecelRamp = 0.1

to_start.InRange = 10

to_start.Straight = False

to_start.Accel = 50

to_start.Decel = 50

to_start.Speed = 50

Robot.Attached = Rb						' Get control of the robot

AC_InitAutoCenter(Rb)					' Initialize global variables

While True

Move.Loc(start_j, to_start)			' Move above stack

Move.WaitForEOM

exc1 = AC_HeightDetect(Rb, 2, -400, -15, stack_top)

If (exc1.ErrorCode >= 0) Then

Console.WriteLine("Detected height:  " & Format(stack_top.Z,"0.00"))

Else

Console.WriteLine(CRLF & "Height detection error: " _

& AC_ErrorMessage(exc1))

Return

End If

End While

Robot.Attached = 0

Return

End Sub

This concludes the description of the HeightDetect function.

Appendix A:  GPL Routine Calling Sequences

Public Sub AC_InitAutoCenter(Rb As Integer)

' ABSTRACT: This is a TOP-LEVEL CALLABLE ROUTINE that must be called

'		before any of the other Auto Center routines are executed.  This

'		routine computes key constants that are utilized by the other

'		routines.  This is done for computational efficiency and clarity.

'		This routine only needs to be executed once, whereas the other

'		Auto Center routines can be executed any number of times.

'

'		This routine should be re-executed if the tool dimensions or any

'		of the robot's link dimensions are modified.

'

' INPUTS:	Rb			Robot number (1-N_ROB).

'

' OUTPUTS:	None

Public Function AC_ErrorMessage(Exc1 As Exception) As String

' ABSTRACT: This is a TOP-LEVEL CALLABLE ROUTINE that converts an Exception

'		object to a string message.  The message will be of the form:

'

'				-error_code, *Error message*

'

'		If the error is not a standard system exception but one generated by

'		this application, it's error code will range from -2000 to -2999.

'		For such errors, this function returns an application specific

'		error message.

'

' INPUTS:	Exc1		Exception to be converted to a string

'

' OUTPUTS:	None

'

' RETURNS:	String that contains the converted error message.

Public Function AC_HeightDetect(Rb As Integer, QuickMode As Integer, _

rXSearchLimit As Double, XMaxForce As Double, _

ByRef NestPosC As Location) As Exception

' ABSTRACT: This is a TOP-LEVEL CALLABLE ROUTINE that detects the height

'		of a stack of plates (or any other objects) below the gripper.

'		When this routine is called, the gripper should be positioned

'		directly above the stack to be tested such that a pure World -Z

'		motion will bring the robot into contact with the stack.  The

'		robot can be holding a plate or the gripper can be empty.

'

'		For best results, the gripper must normally move down at least

'		10 mm before it contacts the surface.  This provides the system

'		with time to accelerate the gripper and obtain a good "free air"

'		reading.

'

'		This routine has been optimized for running with the System Speed

'		set to 100% although slower speeds can be used if desired.

'

'		Prior to calling this routine, "AC_InitAutoCenter" must be

'		executed to initialize key globals that are used in the

'		computation of the forces.

'

'		Also, it is assumed that the ROBOT HAS ALREADY BEEN ATTACHED BY

'		THE CALLING PROGRAM

'

'		At the conclusion of this routine, the gripper is left positioned

'		a few mm above the stack to prevent the robot from continuing to

'		push down on the stack.

'

' INPUTS:	Rb			Robot number (1-N_ROB).

'			QuickMode	Indicates whether this routine should execute

'						quickly or use a more thorough method:

'						 1:	Just execute quick method, 0.5 mm typical

'							and 1.5 mm worse case repeatability errors.

'							About 2-3x less accurate than the thorough

'							method.

'						 2:	Execute quick followed by thorough methods

'							0.3 mm typical and 0.6 mm worse case

'							repeatability errors.  Adds about 4 sec to

'							the total process.

'			rXSearchLimit  This is the maximum distance that the gripper

'						is permitted to move down relative to the starting

'						position, in mm.  If this limit is exceeded before

'						the specified XMaxForce is achieved, an error

'						is signaled.  This value must be a negative number.

'			XMaxForce	Maximum vertical force to be applied when searching

'						for the stack height, in Nt.  This is normally

'						a negative number and -15 is a reasonable value

'						given that there is noise in the free air readings

'						of about +/- 4 Nt.

'

' OUTPUTS:	NestPosC	Cartesian location that contains the detected Z

'						height of the stack (NestPosC.Z).  The other

'						components of the location are roughly the same

'						as the starting position.

'

' RETURNS:	An Exception object that indicate if any error has occurred.